Modeling and Application-Oriented Optimization of Hybrid Multilayer Ferrite Magnetic Shields for SERF-MEG Magnetometers

# Author Name

Abstract--SERF-based magnetoencephalography requires a low-field magnetic environment that combines adequate shielding factor with controlled shield-induced magnetic-noise contribution. This paper revises the ferrite-only multilayer concept into a hybrid multilayer structure in which flexible ferrite sheets form a low-conductivity base and an additional amorphous or nanocrystalline ribbon provides a high-permeability magnetic path. The ferrite base has an inner radius of 100 mm and an axial length of 200 mm. A four-layer ferrite candidate with 0.15 mm sheets and 0.08 mm radial gaps is selected as a manufacturable validation base under total-thickness, radial-space, flexibility, assembly, shielding-factor, and IntH2 constraints, not as a global optimum.

The FEM model uses directional no-shield references and component-based shielding factors. The validated transverse ferrite base gives SFx = 3.629 and IntH2_total = 4.160e-6. Adding an amorphous outer layer increases SFx to 8.548 and reduces IntH2_total to 6.283e-7, while a nanocrystalline outer layer increases SFx to 25.146 and reduces IntH2_total to 5.333e-8 under the present initial-permeability assumptions. Axial cap simulations show that a single nanocrystalline cap improves SFz from 2.412 to 2.907 but introduces an IntH2 tradeoff. Prototype SF measurement, low-field permeability/loss characterization, compensation-coil nulling, and magnetometer PSD measurement remain required before claiming instrument-level sensitivity improvement.

Index Terms--Amorphous ribbon, ferrite magnetic shield, finite-element method, flexible magnetic shielding, magnetic noise, nanocrystalline ribbon, SERF-MEG, shielding factor.

I.  INTRODUCTION

| U LTRA-SENSITIVE systems, including exchange-relaxation-free | magnetic measurement atomic magnetometers, spin- (SERF) comagnetometers, |
| --- | --- |

SERF-based MEG and other optically pumped magnetometer systems require a low-field and low-noise magnetic environment because femtotesla-level biomagnetic signals can be obscured by residual environmental fields and by magnetic noise generated inside the shield itself [1]-[5]. The shield is therefore not only a passive attenuation component. It must be evaluated by directional shielding factor, residual center field, and material-loss-related magnetic-noise indicators.

Author  affiliation,  funding,   and  correspondence  information   should  be added before submission.

Amorphous and nanocrystalline ribbons are commercially available high-permeability soft-magnetic materials. Compared with conventional permalloy or mu-metal shells, ribbon-based amorphous or nanocrystalline layers can be more compatible with low-cost and flexible shield fabrication. These engineering considerations are important for compact SERF-MEG or OPM-oriented shielding, where material availability, assembly, bending, and cost matter in addition to attenuation.

Ferrite materials are attractive for low-noise shielding because of their high electrical resistivity and reduced eddy-current loss. However, the shielding factor of a thin flexible ferrite cylinder can be limited when the available radial space and total thickness are constrained. A purely metallic high-permeability shield can provide stronger attenuation, but may introduce cost, weight, eddy-current, and magnetic-loss concerns. A hybrid ferrite-metallic soft-magnetic structure is therefore investigated to improve attenuation while retaining the low-noise motivation of the ferrite base.

In the proposed hybrid shield, flexible ferrite sheets form a low-conductivity multilayer base, and an additional amorphous or nanocrystalline ribbon forms a high-permeability outer magnetic path. The ferrite base is responsible for flexible low-conductivity shielding, while the outer ribbon is used to raise transverse shielding factor. The combination is intended as a manufacturable composite magnetic shielding route rather than a ferrite-only design.

The present FEM stage uses representative constant relative permeability values as low-field initial-permeability assumptions. Since the shield is intended to operate near the geomagnetic field without a strong excitation field, the initial permeability can be interpreted as the slope of the B-H curve around H = 0. The amorphous and nanocrystalline permeability and loss parameters must be updated after material characterization, and the simulated IntH2 trends must be verified by prototype PSD measurements.

The present study therefore evaluates a hybrid multilayer ferrite shield through a validation-gated evidence chain: a four-layer ferrite base is selected as a manufacturable candidate; amorphous and nanocrystalline outer layers are compared for transverse shielding enhancement; a single-ended cap is evaluated for axial leakage; residual fields are estimated under 50 uT linear scaling; and a P0/P1/P2 prototype route is defined for later SF, permeability, compensation-coil, and PSD validation.

TABLE I

RELATIONSHIP TO REPRESENTATIVE FERRITE SHIELD STUDIES

| Work | Geometry | Main quantity | Scope relative to this work |
| --- | --- | --- | --- |
| Kornack et al. [12] | Single ferrite cylinder | SF and measured noise | No multilayer thin-shell comparison |
| Lee and Romalis [9] | Simple analytic shapes | Magnetic-noise formula | No finite-length multilayer FEM geometry |
| Ma et al. [13] | Cylindrical ferrite shield | SF and IntH2 | Layer-resolved IntH2 not reported |
| Lu et al. [14], [17] | Multi-annular/gapped ferrite shields | Noise-related FEM analysis | Directional no-shield reference not emphasized |
| Sun et al. [16] | Cubic ferrite shield | SF and measured noise | Cubic geometry, not cylindrical multilayer shell |
| This work | Multilayer thin-ferrite cylinder; prior representative ferrite-shield studies are mainly single-layer or monolithic | Directional SF, Vf, layer- wise IntH2 | Measurement-oriented characterization method with validation studies |

This work combines: 1) validated no-shield directional references for component-based shielding-factor evaluation; 2) a virtual pickup-coil model for the geometry-dependent magnetic-noise-related loss integral; 3)

layer-resolved IntH2 decomposition for multilayer ferrite shields; 4) fixed-volume comparison between single-layer and multilayer candidates; 5) unit-consistent auxiliary metrics for low-noise shield design; and 6) validation-gated reporting so that missing axial, segmented, permeability, or hybrid amorphous/nanocrystalline-layer studies are not converted into conclusions.

II. PRINCIPLE OF HYBRID MULTILAYER MAGNETIC SHIELDING

A. Shielding Effectiveness and Magnetic-Noise Coupling

A.   Coordinate System and Measurement Configuration

A multilayer ferrite shield is modeled as a device under test placed at the center of a calibrated magnetic-field region. The cylinder axis is defined as the Cartesian z-axis, while x and y span the transverse plane. Because the shield is a finite open-ended cylinder, transverse and axial shielding are not equivalent. The quantities SFx and SFy describe transverse shielding through the sidewall, whereas SFz describes axial shielding affected by leakage through the open ends [6], [27]–[29]. The same geometry also determines the magnetic-noise coupling because material loss is weighted by the virtual pickup-coil field inside each ferrite layer.

Two finite-element models are used to separate these effects, as shown schematically in Fig. 1. The external uniform-field model gives B0, Bcenter, SFq, and LeakageRatioq and therefore measures shielding effectiveness. The virtual pickup-coil model gives IntH2total and IntH2Li and therefore measures the geometry-dependent loss channel. Multilayer suppression is evaluated by requiring both a smaller directional leakage field and a lower loss indicator under identical material parameters.

For the hybrid structure, the IntH2 integration includes both the ferrite sheets and the added amorphous or nanocrystalline outer layer. The total geometry-dependent indicator is separated as IntH2_total = IntH2_ferrite + IntH2_metal so that the material loss contribution can be discussed without assuming identical imaginary permeability for all materials.

A material-separated noise proxy is therefore proportional to mu''_ferrite IntH2_ferrite + mu''_metal IntH2_metal. A lower IntH2_total is favorable under equal-loss assumptions, but it is not a direct proof of lower measured magnetic noise unless mu''(f), temperature, virtual-coil normalization, and magnetometer PSD data are also available.

The amorphous and nanocrystalline ribbons are modeled by representative initial relative permeability values in the present low-field FEM stage. These values should be read as engineering assumptions for H approximately 0 operation and will be replaced by measured low-field B-H and loss data before final prototype claims.

B.  Directional Shielding Factor

For an external field applied along q  ∈ {x, y, z}, the no- shield reference field at the observation point is

B0  = (B0,x, B0,y, B0,z),                     (1)

and the shielded center leakage field is

Bcenter  = (Bcenter,x, Bcenter,y , Bcenter,z).           (2)

The directional shielding factor is defined by

and the corresponding leakage ratio is

Fig.   1.      Measurement-oriented   characterization   principle.   The   external uniform-field problem provides B0 , Bcenter , SFq, and LeakageRatioq ; the virtual pickup-coil problem provides IntH2 and layer-resolved IntH2.

The  formal  shielding  factor  is  computed  from  the  intended directional component. The vector magnitude BMag  is retained only as an auxiliary consistency check and is not used as the primary SF definition.

C.  No-Shield Directional Reference

The  no-shield  model  is  the  calibration  reference  for  the external-field problem. The ferrite bodies  are deleted  or  as- signed  vacuum  while  the  air  region,  boundary  dimensions, observation  point,   solver  setting,  and  excitation   amplitude are  kept  identical  to  the   shielded  model.  In  vacuum,  for H0  = 1A/m,

B0  = µ0 H0  = (4π × 10-7)(1) = 1.256637 × 10-6  T.   (5)

Thus,  an  ideal  x-directed  reference  should  give  B0,x    ≈ 1.256637  ×  10-6 T,  with  B0,y    and  B0,z    much   smaller. The residual transverse components are treated as numerical residuals caused by finite-element discretization, finite-domain boundary approximation, and interpolation at the observation point.

The reference field is accepted only if the intended compo- nent dominates the other components:

In the post-processing rule set, a failed directional reference prevents the corresponding shielding factors from being used in figures or conclusions.

D.  Analytical Scaling Motivation

The  finite-element  model  is  the  primary  evaluation  tool because the  shield is finite in length, open ended,  and may include layer gaps or segmented slots. Nevertheless, the choice of  design   variables  is  guided  by  the  classical  transverse cylindrical-shell solution. In a source-free region with linear isotropic permeability, ▽ × H = 0 and ▽ · B = 0 allow the magnetic scalar potential to satisfy Laplace’s equation. For an infinitely long coaxial cylindrical shell in a transverse uniform field, the potential in each annular region can be written as

where the constants are obtained from the continuity of normal B and tangential H at each interface  [6]. For a single high- permeability thin shell, this solution gives the familiar leading transverse scaling

showing that  shell thickness, mean radius,  and permeability set the first-order shielding contribution.

For a multilayer coaxial shell, the analytical expansion con- tains direct layer terms and interlayer radius-coupling terms. A typical coupling factor between two layers with mean radii Ri  < Rj  is

This motivates comparing layer number, radial gap, and total radial occupation rather than comparing total thickness alone. The present manuscript uses this analytical structure only as a scaling motivation; the reported SFq  values are obtained from the validated directional FEM reference in (3), not from the ideal infinite-cylinder expression.

Under the thin-shell  approximation,  a  set  of equal-height coaxial layers with total ferrite thickness Tf  = NLa has

Vf  ≈ 2πLRcTf ,                            (10)

where Rc  is the radial-window center radius. Therefore, fixed Tf  is a useful first-order proxy for fixed ferrite volume. In the actual post-processing, however, Vf  is always computed from the exact layer radii in (22) or from CAD volume export.

E.  Virtual Pickup-Coil Model and IntH2

The shield-induced magnetic noise can be related to material loss through the fluctuation-dissipation theorem  [9],  [18].  A virtual pickup coil placed at the observation point produces a magnetic field Hvc in the ferrite. For coil area A, turn number Nc ,  peak  current  I0 ,  and  magnetic  moment  m  =  ANc I0 , the  hysteresis-loss  contribution  to  the  magnetic-field  noise amplitude spectral density can be written as

where µ′′  = µ0 µ is the absolute imaginary permeability if

µ is specified as a relative quantity. The current I0  is a peak

current; if an RMS current is used in another implementation, the normalization must be converted consistently.

This manuscript does not claim measured magnetic-noise spectra. The exported FEM quantity is the geometry-dependent magnetic-noise-related loss indicator

with SI unit A2 m. IntH2 becomes an absolute noise predic-

tion only after measured µ(f), temperature, frequency, and

virtual-coil normalization  are  specified. Relative comparison by IntH2 is meaningful only for identical material properties, identical coil normalization, and identical observation direc- tion.

For  two  candidate  geometries  A  and  B  evaluated  under identical  material  loss   and  virtual-coil  normalization,   (11)

gives the relative loss-induced field-noise scaling

Thus,  IntH2  ratios  should  not  be  interpreted  linearly  as magnetic-noise  ratios;  the  square-root  relation  applies  only under the identical-normalization assumptions stated above.

For a multilayer shield, each ferrite layer is retained as an independent body. The layer-resolved terms are

and

III. FEM MODEL AND HYBRID STRUCTURAL DESIGN MATRIX

The export is accepted only if

III. MEASUREMENT AND CALCULATION OF MAGNETIC SHIELDING UNDER MULTILAYER STRUCTURE

A. Multilayer Structure and Mechanical Feasibility

The DUT is an open-ended multilayer cylindrical shell assembled conceptually from flexible-composite ferrite sheets. The mechanical objective is to avoid a brittle monolithic large-diameter ferrite cylinder while retaining a cylindrical magnetic path. The present manuscript does not report prototype mechanical test data; instead, it defines the manufacturable geometry that will be used for later assembly tests. The model has inner radius Rin, axial length L, layer count NL, ferrite thickness a, and radial interlayer gap g. The current design matrix uses

Rin  = 100mm,        L = 200mm.              (17)

The total ferrite thickness and radial occupation are

Tf  = NLa,                                (18)

Tspace  = NLa + (NL  — 1)g.                  (19)

For layer i, where i = 1 is the innermost layer,

Ri,in  = Rin + (i — 1)(a + g),                (20)

Ri,out  = Ri,in + a.                                   (21)

The ferrite volume is  calculated from the  exact  cylindrical- shell expression

## Fig. 2.   Geometry definition for the multilayer thin-ferrite cylindrical shell.

Although Vf  is displayed in mm3  for engineering readability, all SI-derived metrics are computed after unit conversion.

B. FEM Calculation Matrix and Data Export

The calculation matrix is expanded from ferrite-only transverse continuous shells to a hybrid evidence chain. It includes the four-layer ferrite base, ferrite plus amorphous outer layer, ferrite plus nanocrystalline outer layer, a screened P0 amorphous shell-cap reference, and a four-layer ferrite case with a single nanocrystalline axial cap. The P0 case is intentionally not a no-shield reference, because the physical SERF residual field may be too large to measure without shielding.

Each external-field case exports Bcenter components and the appropriate no-shield B0 reference. Each virtual pickup-coil case exports IntH2_total and material- or layer-resolved contributions, including IntH2_outer for hybrid outer layers. Missing values are marked as missing or pending and are not replaced by zero.

B. FEM Calculation Matrix and Data Export

The present main matrix is a transverse continuous-shell calculation study. It is intentionally separated from prototype measurement, axial, segmented, permeability-sensitivity, and convergence studies so that each validation stage can be tracked independently. Table II lists the current main cases.

Each    shielded    external-field    case    exports    Bcenter,x , Bcenter,y , Bcenter,z, and Bcenter ,Mag. The no-shield reference exports B0,x, B0,y, B0,z, and B0 ,Mag. Each virtual pickup-coil case exports IntH2total  and the layerwise quantities IntH2Li . The post-processing scripts merge raw CSV exports by case ID and field direction, compute the derived metrics, and mark missing or failed rows without substituting zero values.

C.  Comparison Metrics

The volume-normalized shielding index is defined as

It is an auxiliary index introduced for this comparison, not a material property. When Vf  is reported in mm3 , the displayed unit is mm-3; SI-derived quantities are computed after con- verting volume to m3.

A dimensionless normalized form is

where the reference is the continuous single-layer shell with a  =  0.60mm.  The  loss-integral  density  and  attenuation- normalized loss indicator are

The metrics are interpreted jointly with Tspace  and the vali- dation status. No single auxiliary metric defines a final shield geometry.

E. Prototype Measurement and Mechanical-Test Route

No prototype shielding-factor or magnetometer PSD result is claimed in the present manuscript. The planned validation route uses three practical configurations: P0, a single amorphous shell with a single +Z amorphous cap; P1, a four-layer ferrite prototype; and P2, a four-layer ferrite prototype with the selected nanocrystalline outer layer. Optional two-layer ferrite samples may be added to document bending and assembly limits.

IV. SIMULATION RESULTS AND STRUCTURAL OPTIMIZATION

A. Four-Layer Ferrite Base and Hybrid Outer-Layer Enhancement

The x-directed no-shield reference gives B0,x = 1.256637e-6 T and passes the direction-purity check. The four-layer ferrite base C2_N4_t015_g008_x gives SFx = 3.629 and IntH2_total = 4.160e-6. This confirms the base as a manufacturable low-conductivity candidate, but the attenuation is limited if the ferrite stack is used alone. Compared with this ferrite-only base, the amorphous outer layer increases SFx to 8.548, while the nanocrystalline outer layer increases SFx to 25.146. Meanwhile, IntH2_total decreases to 6.283e-7 and 5.333e-8 for the amorphous and nanocrystalline hybrid cases, respectively. The IntH2 integration includes both the ferrite layers and the added outer high-permeability layer.

B. Material-Separated Magnetic-Noise-Related Loss

The decrease in IntH2_total indicates a favorable geometry-dependent trend under the current material assumptions, but it does not by itself prove that measured shield noise will decrease. For hybrid structures, the relevant noise proxy depends on both ferrite and metallic-ribbon loss: mu''_ferrite IntH2_ferrite + mu''_metal IntH2_metal. The nanocrystalline outer layer is therefore identified as the current simulated priority candidate because it gives the highest SFx and the lowest IntH2_total, while final material-noise claims remain conditional on measured mu'' and PSD validation.

C. Axial Single-Cap Tradeoff and Geomagnetic Residual Estimate

For axial shielding, the open four-layer ferrite cylinder gives SFz = 2.412 and LeakageRatioz approximately 0.4145. The P0 amorphous shell-cap reference gives SFz = 2.526 and IntH2_total = 3.960e-7. A four-layer ferrite shield with a single +Z nanocrystalline cap improves SFz to 2.907 and reduces the leakage ratio to 0.3440, but the total IntH2 increases relative to the open ferrite axial case. This is treated as an SFz-IntH2 tradeoff, not as a demonstrated magnetometer-sensitivity improvement. In the axial cap simulation, a small axial gap is introduced between the cylinder and cap to represent mechanical support and to avoid an unrealistically sharp zero-gap FEM contact.

C. Practical Application in SERF-MEG Shielding

The cases C1_N1_t060_x, C2_N3_t020_g010_x, and C2_N4_t015_g008_x have nearly equal ferrite volumes and provide a first application-oriented comparison for a compact SERF-MEG shield. Among these validated transverse continuous-shell candidates, the four-layer case has the largest SFx and the smallest χH. This indicates a better simulated tradeoff between leakage-field attenuation and loss-related coupling, but it is not yet an instrument sensitivity result. A SERF magnetometer must still be placed inside the assembled multilayer shield to measure residual field stability and sensitivity before an application-level specification can be reported.

TABLE II

MAIN CONTINUOUS-SHELL TRANSVERSE MATRIX

| Case | NL | a (mm) | g (mm) | Direction | Purpose |
| --- | --- | --- | --- | --- | --- |
| B0 reference x | 0 | 0 | 0 | x | no-shield reference |
| C1 N1 t008 x | 1 | 0.08 | 0 | x | thinnest single layer |
| C1 N1 t020 x | 1 | 0.20 | 0 | x | thin single layer |
| C1 N1 t040 x | 1 | 0.40 | 0 | x | intermediate single layer |
| C1 N1 t060 x | 1 | 0.60 | 0 | x | equal-volume single-layer baseline |
| C2 N2 t020 g010 x | 2 | 0.20 | 0.10 | x | multilayer candidate |
| C2 N3 t020 g010 x | 3 | 0.20 | 0.10 | x | equal-volume multilayer candidate |
| C2 N4 t015 g008 x | 4 | 0.15 | 0.08 | x | equal-volume multilayer candidate |
| H_C2_N4_outer_AM | 4 + AM | 0.15 + 0.20 | 0.08/0.40 | x | hybrid amorphous outer-layer case |
| H_C2_N4_outer_NC | 4 + NC | 0.15 + 0.20 | 0.08/0.40 | x | hybrid nanocrystalline outer-layer case |
| P0_AM_shell_cap | AM shell + cap | 0.20 | asm. | z | screened prototype reference, not no-shield |
| ZCAP_C2_N4_NC | 4 + NC cap | 0.15 + 0.20 | axial | z | axial single-cap tradeoff case |

TABLE III

NO-SHIELD DIRECTIONAL REFERENCE VALIDATION

| Reference | Dir. | ¦Bx¦ (T) | ¦By¦ (T) | ¦Bz¦ (T) | Status |
| --- | --- | --- | --- | --- | --- |
| B0 reference x | x | 1.257e-6 | 3.828e-15 | 1.943e-13 | passed |
| B0 reference z | z | 2.468e-15 | 3.361e-14 | 1.257e-6 | passed |

## Fig. 3.   Transverse continuous-shell results: SFx  versus ferrite volume.

## Fig. 4.  Equal-volume transverse comparison using SFx, IntH2, η, and χH .

D.  Layer-Resolved IntH2 Distribution

Layerwise IntH2 summation validation passed for all seven transverse continuous-shell cases. The multilayer fractions are listed in Table VI and plotted in Fig. 5. The four-layer case shows the largest fraction in the outermost layer under the present virtual-coil normalization. This is interpreted as a geometry-dependent integral distribution, not as a measured magnetic-noise spectrum.

E. Mesh, Boundary-Domain, and Model-Use Caveats

Mesh convergence data remain available for representative ferrite cases. These checks support scalar trend reporting but not a complete numerical-robustness claim because total element count, adaptive-pass count, and air-domain metadata are incomplete. Hybrid and cap results are therefore reported as current FEM evidence requiring continued numerical and prototype validation.

F. Axial Shielding Characterization

The z-directed no-shield reference has passed direction validation. Current axial results are reported for the open four-layer ferrite cylinder, the P0 amorphous shell-cap reference, and the four-layer ferrite case with a single nanocrystalline cap. These results are used only to describe an axial SFz-IntH2 tradeoff and are not generalized to a final closed-shield design.

G. Segmented-Shell and Assembly Limits

Segmented aligned and staggered models are still required to quantify manufacturable slot penalties. The continuous-shell results should therefore be read as idealized magnetic-shell evidence. The physical prototype plan must still check bending stiffness, assembly gaps, and whether four ferrite sheets are near the practical mounting limit.

For the amorphous and nanocrystalline additions, the relative permeability values used here represent initial-permeability assumptions for low-field operation. They are placeholders for the slope of the measured B-H curve around H = 0 and must be updated after material characterization.

The application-oriented interpretation is likewise conservative. The results suggest a potential route toward lightweight OPM/SERF shielding, but a practical system still requires P0/P1/P2 SF measurements, compensation-coil nulling, and magnetometer PSD validation.

TABLE IV

VALIDATED TRANSVERSE CONTINUOUS-SHELL RESULTS

| Case | NL | a (mm) | g Vf (mm) (mm3) | Tspace (mm) | Bcenter,x (T) | SFx | Leakage ratio | IntH2 (A2m) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| C1 N1 t008 x | 1 | 0.08 | 0.00 10057.1 | 0.08 | 9.375e-7 | 1.340 | 0.746 | 3.819e-6 |
| C1 N1 t020 x | 1 | 0.20 | 0.00 25157.9 | 0.20 | 8.353e-7 | 1.504 | 0.665 | 4.383e-6 |
| C1 N1 t040 x | 1 | 0.40 | 0.00 50366.0 | 0.40 | 5.832e-7 | 2.155 | 0.464 | 4.973e-6 |
| C1 N1 t060 x | 1 | 0.60 | 0.00 75624.4 | 0.60 | 4.418e-7 | 2.845 | 0.352 | 4.649e-6 |
| C2 N2 t020 g010 x | 2 | 0.20 | 0.10 50391.1 | 0.50 | 5.616e-7 | 2.238 | 0.447 | 4.917e-6 |
| C2 N3 t020 g010 x | 3 | 0.20 | 0.10 75699.8 | 0.80 | 4.197e-7 | 2.994 | 0.334 | 4.381e-6 |
| C2 N4 t015 g008 x | 4 | 0.15 | 0.08 75714.9 | 0.84 | 3.463e-7 | 3.629 | 0.276 | 4.160e-6 |

# TABLE V

EQUAL-VOLUME TRANSVERSE COMPARISON

| Case | Vf (mm3) | SFx | IntH2 (A2m) | η | χH |
| --- | --- | --- | --- | --- | --- |
| C1 N1 t060 x | 75624.4 | 2.845 | 4.649e-6 | 1.045 | 4.447e-6 |
| C2 N3 t020 g010 x | 75699.8 | 2.994 | 4.381e-6 | 1.095 | 3.995e-6 |
| C2 N4 t015 g008 x | 75714.9 | 3.629 | 4.160e-6 | 1.287 | 3.227e-6 |

# TABLE VI

LAYERWISE INTH2 FRACTIONS FOR MULTILAYER CASES

| Case | L1 (%) | L2 (%) | L3 (%) | L4 (%) |
| --- | --- | --- | --- | --- |
| C2 N2 t020 g010 x | 49.4 | 50.6 | — | — |
| C2 N3 t020 g010 x | 25.7 | 36.1 | 38.2 | — |
| C2 N4 t015 g008 x | 19.5 | 19.8 | 20.6 | 40.0 |

## Fig. 6.   Mesh convergence of SFx  for the thinnest single-layer case and the four-layer candidate. Mesh metadata remain incomplete.

H. Initial-Permeability and Material-Loss Sensitivity

The amorphous and nanocrystalline ribbons are modeled with representative initial relative permeability values for low-field operation. These constants are not arbitrary final material claims; they approximate the low-field B-H slope around H = 0 under geomagnetic-field operation. The model must be updated after measured permeability and magnetic-loss data are obtained.

C2  N4  t015  g008  x  at  µ  = 500,  1000,  2000, and 5000 is

I. Application-Oriented Interpretation

Under linear scaling to a 50 uT external geomagnetic field, the estimated transverse residual field decreases from 13.78 uT for the ferrite-only base to 5.85 uT with the amorphous outer layer and 1.99 uT with the nanocrystalline outer layer. These values are linear initial-permeability estimates, not nonlinear geomagnetic B-H simulations.

The long-term application target is a lightweight and potentially wearable magnetic shielding route for OPM- or SERF-based biomagnetic measurements. The present work supports this direction through modeling and prototype planning, but does not claim that wearable MEG or final SERF sensitivity has already been realized.

flux  lines,  Hvc ,  and  Hc   contours   are  not  yet  available.

The revised results show that the ferrite-only multilayer base is useful as a low-conductivity manufacturable platform, but hybridization is needed to raise the shielding factor toward an application-oriented range. The nanocrystalline outer layer gives the strongest simulated transverse attenuation and the lowest IntH2_total in the current dataset. This preference is conditional on the initial-permeability assumptions and must be revisited after measured permeability and loss data are available.

The four-layer ferrite base is not claimed as a global optimum.

Although additional ferrite sheets may further improve shielding, the multilayer stack becomes increasingly difficult to bend and assemble.

The four-layer ferrite base is therefore selected as a manufacturable validation candidate under the present total-thickness, radial-space, flexibility, and assembly constraints. The P0/P1/P2 prototype plan is designed to test this engineering claim rather than to replace it with unsupported simulation-only conclusions.

The IntH2 results should be read as magnetic-noise-related loss indicators, not measured magnetic noise. For hybrid structures, material loss must be separated into ferrite and metallic-ribbon terms. A lower IntH2_total under equal-loss assumptions is favorable, but the final sensitivity claim requires prototype SF measurements, measured mu''(f), compensation-coil nulling, and magnetometer PSD data.

TABLE VII

MESH CONVERGENCE FOR THE REPRESENTATIVE TRANSVERSE CASES

| Case | Mesh level | Total elements | Across thickness | Adaptive passes | Air scale | Bcenter,x (T) | SFx | IntH2 (A2m) | Rel. change (SFx/IntH2) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| C1 N1 t008 x | coarse | — | 1 | — | — | 9.265e-7 | 1.356 | 3.800e-6 | —/— |
| C1 N1 t008 x | medium | — | 3 | — | — | 9.375e-7 | 1.340 | 3.819e-6 | 1.19e-2/5.00e-3 |
| C1 N1 t008 x | fine | — | 5 | — | — | 9.375e-7 | 1.340 | 3.819e-6 | 1.72e-10/1.81e-10 |
| C2 N4 t015 g008 x | coarse | — | 1 | — | — | 3.463e-7 | 3.629 | 4.160e-6 | —/— |
| C2 N4 t015 g008 x | medium | — | 3 | — | — | 3.463e-7 | 3.629 | 4.160e-6 | 3.42e-10/8.17e-11 |
| C2 N4 t015 g008 x | fine | — | 5 | — | — | 3.463e-7 | 3.629 | 4.160e-6 | 1.93e-10/4.33e-11 |

The application claim is also conservative. Linear scaling to a 50 uT geomagnetic field indicates residual fields at the uT level, so three-axis compensation coils are required before SERF operation. Therefore, the paper presents a potential route toward compact OPM/SERF shielding rather than a completed wearable MEG demonstration.

volume ranking  cannot yet be  called robust to permeability

variation. The next required  sensitivity  sweep at µ   = 500,

A hybrid multilayer ferrite magnetic shield has been formulated for SERF-MEG magnetometer applications. The four-layer ferrite base is selected as a manufacturable validation candidate under thickness, flexibility, assembly, SF, and IntH2 constraints, not as a global optimum. In transverse simulations, the ferrite-only base gives SFx = 3.629 and IntH2_total = 4.160e-6. Adding an amorphous outer layer increases SFx to 8.548, and adding a nanocrystalline outer layer increases SFx to 25.146 while reducing IntH2_total to 5.333e-8 under the present initial-permeability assumptions.

Axial cap simulations show that a single nanocrystalline cap improves SFz but introduces an IntH2 tradeoff. The current conclusions remain conditional: the amorphous and nanocrystalline permeability values represent low-field initial-permeability assumptions, material loss must be measured, and prototype P0/P1/P2 SF and PSD tests are required before claiming instrument-level sensitivity improvement.

A measurement-oriented characterization method for mul- tilayer  thin-ferrite  cylindrical   shields  has  been   established and  applied  to  validated  transverse  continuous-shell  AEDT exports. The method uses a passed no-shield directional ref- erence,  component-based  SFx ,  layer-resolved  IntH2  checks, exact  ferrite-volume  calculation,  and  unit-consistent   auxil- iary metrics. For the  seven  validated transverse  continuous- shell candidates, the equal-volume four-layer design gives the highest  SFx  and the lowest IntH2-normalized auxiliary loss indicators among the candidates evaluated. Mesh convergence checks  for  the  thinnest  single-layer  and  selected  four-layer cases support the reported scalar values, but the mesh metadata are incomplete and boundary-domain convergence has not yet been completed.

The    current    conclusions    are    limited    to    transverse continuous-shell  cases  with   screening  material  parameters. They   do   not   establish    axial   performance,    segmented- assembly    performance,    complete    numerical    robustness, permeability-variation robustness,  or  a  final instrument-level recommendation. IntH2 is treated only as a magnetic-noise- related loss indicator; absolute magnetic-noise spectra require

measured µ(f) and experimental validation.

The prototype validation plan uses P0, a single amorphous shell plus single +Z amorphous cap; P1, a four-layer ferrite prototype; and P2, a four-layer ferrite prototype with a selected nanocrystalline outer layer. The required measurements include x- and z-directed SF, low-field permeability and loss characterization, compensation-coil residual-field nulling, and magnetometer PSD.

APPENDIX A

PROTOTYPE SF MEASUREMENT NOTE

No experimental  shielding-factor data  are included in the present manuscript. A future prototype measurement  should use a calibrated Helmholtz or triaxial coil system and a cal- ibrated magnetic sensor, with B0  and Bcenter  evaluated from

the same directional component used in the FEM definition. For independent measurements,

For correlated measurements,

u(SFq,exp) = u(B0 )+ u(Bcenter )—2ρur (B0 )ur (Bcenter ).

(29)

ACKNOWLEDGMENT

Funding,   facilities,    and   collaborator    acknowledgments should be added before submission.

REFERENCES

[1]  J.  C.  Allred,  R.  N.  Lyman,  T.  W.  Kornack,  and  M.  V.  Romalis, “High-sensitivity  atomic  magnetometer  unaffected  by   spin-exchange relaxation,” Phys. Rev. Lett., vol. 89, no. 13, Sep. 2002, Art. no. 130801.

[2]  I.  K.  Kominis,  T.  W.  Kornack,  J.  C.  Allred,  and  M.  V.  Romalis,  “A subfemtotesla  multichannel  atomic  magnetometer,”  Nature,  vol.  422, no. 6932, pp. 596–599, Apr. 2003.

[3]  M. P. Ledbetter, I. M. Savukov, V. M. Acosta, D. Budker, and M. V. Ro- malis,  “Spin-exchange-relaxation-free  magnetometry  with  Cs  vapor,” Phys. Rev. A, vol. 77, no. 3, Mar. 2008, Art. no. 033408.

[4]  S.  Baillet,  “Magnetoencephalography  for  brain  electrophysiology  and imaging,” Nature Neurosci., vol. 20, no. 3, pp. 327–339, Mar. 2017.

[5]  T.  E.  Chupp,  P.  Fierlinger,  M.  J.  Ramsey-Musolf,  and  J.  T.  Singh, “Electric  dipole moments  of atoms, molecules, nuclei,  and particles,” Rev. Mod. Phys., vol. 91, no. 1, Jan. 2019, Art. no. 015001.

[6]  T. J. Sumner, J. M. Pendlebury, and K. F. Smith, “Conventional magnetic shielding,”  J.  Phys.  D:  Appl.  Phys.,  vol.  20,  no.  9,  pp.   1095–1101, Sep. 1987.

[7]  V.  Kelha,  J.  Pukki,  R.  Peltonen,  A.  Penttinen,  R.  Ilmoniemi,  and J.  Heino,  “Design,  construction,  and  performance  of  a  large-volume magnetic  shield,”  IEEE  Trans.  Magn.,  vol.   18,  no.   1,  pp.  260–270, Jan. 1982.

[8]  I. Altarev et al., “A magnetically shielded room with ultra low residual field  and  gradient,”  Rev.  Sci.  Instrum.,  vol.  85,  no.  7,  Jul.  2014, Art. no. 075106.

[9]  S.-K. Lee and M. V. Romalis, “Calculation of magnetic field noise from high-permeability magnetic shields and conducting objects with simple geometry,” J. Appl. Phys., vol.  103, no. 8, Apr. 2008, Art. no. 084904.

[10]  H. B. Dang, A. C. Maloof, and M. V. Romalis, “Ultrahigh sensitivity magnetic field and magnetization measurements with an atomic magne- tometer,” Appl. Phys. Lett., vol. 97, no. 15, Oct. 2010, Art. no. 151110.

[11]  D. Ma et al., “A novel low-noise mu-metal magnetic shield with winding shape,” Sens. Actuators A, Phys., vol. 346, Oct. 2022, Art. no.  113884.

[12]  T.  W.  Kornack,   S.  J.   Smullin,  S.-K.  Lee,   and  M.  V.  Romalis,   “A low-noise  ferrite magnetic  shield,” Appl. Phys. Lett.,  vol.  90, no.  22, May 2007, Art. no. 223501.

[13]  D.  Ma  et  al.,  “Parameter  modeling  analysis  of  a  cylindrical  ferrite magnetic shield to reduce magnetic noise,” IEEE Trans. Ind. Electron., vol. 69, no. 1, pp. 991–998, Jan. 2022.

[14]  J. Lu et al., “Study of magnetic noise of a multi-annular ferrite shield,” IEEE Access, vol. 8, pp. 40918–40924, 2020.

[15]  B.  Sun  et al.,  “Correlating  the  microstructure  of Mn–Zn  ferrite  with magnetic noise for magnetic shield applications,” Ceram. Int., vol. 49, no. 8, pp. 11960–11967, Apr. 2023.

[16]  B. Sun, D. Ma, X. Fang, Y. Xue, J. Lu, H. Chen, M. Zhang, H. Wei, B. Han, and Y. Zhai, “Suppression of magnetic noise and field in cubic low-noise ferrite magnetic shields,” IEEE Trans. Instrum. Meas., vol. 73, pp. 1–10, 2024, Art. no. 1501810.

[17]  J.  Lu  et  al.,  “Effect  of  gaps  on  magnetic  noise  of  cylindrical  fer- rite  shield,”  J.  Phys.  D:  Appl.  Phys.,   vol.  54,  no.   25,  Jun.   2021, Art. no. 255002.

[18]  R.  Kubo,  “The  fluctuation-dissipation  theorem,”  Rep.  Prog.  Phys., vol. 29, no. 1, pp. 255–284, Jan. 1966.

[19]  C.  Liu  et al.,  “Investigation  on  the  effects  of micro-vibration  on  the atomic comagnetometer,” IEEE  Trans. Instrum. Meas., vol. 72, 2023, Art. no. 9511711.

[20]  P. Dutta  and P. M. Horn,  “Low-frequency  fluctuations in  solids:  1/f noise,” Rev. Mod. Phys., vol. 53, no. 3, pp. 497–516, Jul. 1981.

[21]  K. Yang et al., “Improved measurement of the low-frequency complex permeability of ferrite annulus for low-noise magnetic shielding,” IEEE Access, vol. 7, pp. 126059–126065, 2019.

[22]  K.  Yang  et al.,  “Minimizing  magnetic  fields  of the  low-noise  MnZn ferrite  magnetic  shield  for  atomic  magnetometer,”  J.  Phys.  D:  Appl. Phys., vol. 55, no. 1, Jan. 2022, Art. no. 015003.

[23]  Z. Xu, Z. Zhang, B. Wu, H. Wang, X. Kong, and M. Wang, “A study of enclosed magnetic shielding room by simulation,” IEEE Trans. Appl. Supercond., vol. 31, no. 8, pp. 1–5, Nov. 2021.

[24]  X. Xu, L. Wang, W. Liu, and Z. Zhao, “Theoretical modeling and char- acterization  of equivalent magnetic properties in laminated composite magnetic shielding,” Measurement, vol. 251, 2025, Art. no. 117237.

[25]   Cores Made of Soft Magnetic Materials—Measuring Methods—Part 2: Magnetic Properties at Low Excitation Level, IEC 62044-2, 2005.

[26]  TDK  Corporation,  “Mn–Zn  ferrite  material  characteristics:  PC95  and related  Mn–Zn  ferrite  grades,”  product  catalog,  accessed  May  2026. [Online].      Available:      https://product.tdk.com/system/files/dam/doc/ product/ferrite/ferrite/ferrite-core/catalog/ferrite  mn-zn  material

characteristics  en.pdf

[27]  A. Mager, “Magnetic shields,” IEEE Trans. Magn., vol. 6, no. 1, pp. 67– 75, Mar. 1970.

[28]  K. Nagashima, I. Sasada, and K. Tashiro, “High-performance bench-top cylindrical magnetic shield with magnetic shaking enhancement,” IEEE Trans. Magn., vol. 38, no. 5, pp. 3335–3337, Sep. 2002.

[29]  J. M. G. D. K. O’Connell, J. M. Brown, and T. W. Kornack, “Technique for high axial  shielding factor performance of large-scale, thin, open- ended, cylindrical Metglas magnetic shields,” arXiv:1107.2625, 2011.

[30]  J.  Fang,  S.  Wan,  J.  Qin,  and  Y.  Chen,  “A  novel  method  for  in-situ measuring the static shielding factor of a magnetic shield,” Measurement, vol. 170, 2021, Art. no. 108718.

[31]  T. Brys  et al.,  “Magnetic  field  stabilization for magnetically  shielded volumes by external field coils,” J. Appl. Phys., vol.  116, no. 8, 2014, Art. no. 084903.